# Internal Engineering: Hardware Construction Plan for Active and Pro Smart Ring

## Design Basis and Non-Negotiable Architecture Principles

This plan uses the supplied product specification that defines Active as the fitness-focused, motion-optimized version and Pro as the all-day wellness, sleep, stress, and advanced-insight version. Both versions must share the same validated core architecture for HR, HRV, resting HR, sleep, activity, and temperature baseline; Pro may add redundancy, higher sampling flexibility, and advanced sensors, but Active must not be treated as a lower-accuracy product.

The technical architecture also follows the background sensor-analysis principle that ring accuracy is not created by one sensor alone. It comes from the full stack: optical layout, analog front-end design, flex-PCB miniaturization, mechanical fit, power management, firmware filtering, calibration, and signal-quality algorithms.

The clean-sheet goal is therefore:

One ring family, two personalities:

- Active: thinner, tougher, sweat-proof, motion-stable, gym-safe.

- Pro: thinner but more premium, sleep/resting-accuracy optimized, redundant sensing, advanced stress/temperature/sleep analytics.

# 1. Common Core Platform for Both Versions

## 1.1 Shared Ring Stack

Both versions use a three-layer radial construction:

- Outer fashion shell

  - Structural titanium or ceramic.

  - Seamless decorative exterior.

  - No external sensor bump, pod, ridge, camera hole, or plastic module.

- Middle electronics belt

  - Ultra-thin curved rigid-flex PCB.

  - Curved thin-film battery segments.

  - Optical AFE, MCU, IMU, wireless, PMIC.

  - Black optical-isolation ribs molded directly into the inner carrier.

- Inner biometric interface

  - Smooth hypoallergenic liner.

  - Flush optical windows.

  - Thermal islands for skin-temperature sensing.

  - Compliant contact rails that improve skin contact without visible bumps.

# 2. Active Version Hardware Plan

## 2.1 Product Role

Active is the motion-optimized version for fitness-focused users. It prioritizes:

- Reliable activity tracking.

- Sweat and water resistance.

- Robust sensor readings during movement.

- A thinner, fashionable ring form.

The Product Idea source defines Active around trustworthy sleep, HR, HRV, activity, and temperature essentials, with multi-wavelength PPG, skin temperature, a 3-axis accelerometer baseline, smart sampling, durability, and 3.0–4.0 g weight target.

## 2.2 Active Sensor Subsystem Architecture

### Biometric Sensors

1. Reflectance PPG module

- Type: multi-wavelength optical PPG.

- Wavelengths:

  - Green: HR during movement.

  - Red: SpO₂ support.

  - Infrared: sleep/resting HR, HRV, vascular waveform stability.

- Package: side-firing micro-LED dies plus low-profile photodiodes.

- Use cases:

  - HR.

  - Resting HR.

  - HRV.

  - SpO₂ trend.

  - Pulse waveform features.

2. Skin-temperature sensor

- Type: fast-response NTC thermistor or silicon temperature IC thermally coupled to inner wall.

- Placement: inner-ring thermal island at palmar-side contact zone.

- Use cases:

  - Daily temperature baseline.

  - Rapid thermal deviation tracking.

Fever/illness trend support.

3. Motion sensor

- Type: ultra-low-power 6-axis IMU.

- Normal mode: accelerometer active, gyroscope duty-cycled.

- Exercise mode: accelerometer + gyroscope enabled.

- Use cases:

  - Steps.

  - Sleep motion.

  - Activity classification.

  - Motion-artifact rejection for PPG.

4. Capacitive skin-contact / fit sensor

- Type: low-power capacitive contact pads around inner liner.

- Use cases:

  - Detect loose ring fit.

  - Reject poor PPG readings.

## 2.3 Active Optical Layout

Target: fit the optical system within ≤2.05 mm radial wall thickness, with no raised internal bumps.

### Layout

- Two side-firing LED packages

  - Each package contains green, red, and IR dies.

  - Mounted tangentially on the flex PCB so light is injected sideways into skin through shallow optical wells.

- Two large-area photodiodes

  - One central photodiode for primary return signal.

  - One offset photodiode for alternate return path and motion-artifact rejection.

- Four primary optical paths

  - LED A → PD 1.

  - LED A → PD 2.

  - LED B → PD 1.

  - LED B → PD 2.

### Mechanical Optical Features

- Flush sapphire or aluminosilicate glass windows, 0.15–0.20 mm thick.

- Black epoxy isolation wall between emitters and photodiodes.

- Recessed optical cavities filled with optically clear biocompatible resin.

- No internal sensor bump; instead, the entire inner wall has a shallow ergonomic concavity.

### Active Rationale

Active uses fewer optical channels than Pro to preserve battery and free board area for reinforced waterproofing and impact durability. The design is optimized for motion stability, not maximum sensor count.

## 2.4 Active Mechanical and Industrial Design

### Target Dimensions

- Width: 5.2–5.8 mm

- Radial wall thickness: 1.80–2.05 mm

- Mass: 3.0–4.0 g

- Outer profile: soft rectangular-oval or shallow concave fashion band

- Sensor protrusions: none

### Materials

Exterior

- Grade 5 titanium with matte ceramic-like PVD coating, or zirconia ceramic edition.

- Finish options:

  - Matte black.

  - Brushed silver.

  - Graphite.

  - Soft gold.

Interior

- Hypoallergenic medical-grade epoxy-polymer liner.

- Local thermal islands under temperature sensor.

- Flush optical windows.

- Sweat-resistant, soap-resistant, UV-stable inner coating.

### Water and Sweat Resistance

- Target: IP69K

- Design choices:

  - Laser-welded or diffusion-bonded outer shell.

  - No open mechanical buttons.

  - Fully potted electronics except battery expansion-relief zone.

  - Salt-sweat corrosion test for gym use.

### Sensor-Skin Contact Without Thickness Growth

- Inner wall uses broad, smooth contact rails rather than bumps.

- Compliant liner deflection: 40–80 µm.

- Flex-PCB includes local omega-loop strain relief behind optical windows.

- Ring sizing kit must bias toward snug-but-comfortable fit because optical performance depends on stable skin contact.

## 2.5 Active Electronics and Processing Unit

### MCU

- Low-power ARM Cortex-M4F or Cortex-M33.

- Required features:

  - Floating-point DSP.

  - BLE stack.

  - Secure boot.

  - Hardware crypto.

  - Multiple low-power sleep states.

- Minimum resources:

  - 1 MB flash.

  - 256 KB RAM.

  - 8–16 MB external low-power flash for event logs.

### Analog Front-End

- One multi-channel optical AFE.

- Integrated LED drivers.

- Programmable transimpedance amplifiers.

- Ambient-light cancellation.

- Low-noise PPG acquisition during motion.

### Battery

- Curved thin-film lithium cell.

- Capacity: 15 mAh

- Charging:

  - Magnetic pogo or inductive charging cradle.

  - No open USB port on ring.

### Wireless

  - BLE for phone and gateway communication.

## 2.6 Active Firmware and Signal Processing

### Normal Sampling

- PPG: low-duty periodic sampling, higher during sleep and post-exercise recovery.

- IMU: continuous low-rate accelerometer, gyro off unless activity/risk detected.

- Temperature: periodic baseline sampling.

- Fit sensing: periodic and before high-confidence readings.

# 3. Pro Version Hardware Plan

## 3.1 Product Role

Pro is the advanced all-day wellness and sleep version. It prioritizes:

- Better resting signal redundancy.

- Higher-resolution HRV.

- Advanced stress and recovery.

- Better sleep tracking.

- More stable temperature trends.

- Premium materials.

The Product Idea source defines Pro as everything in Active plus higher-density PPG, additional temperature points, EDA/stress sensing, optional haptics, more onboard memory, a more advanced motion package, premium titanium or ceramic materials, and 3.5–4.5 g target weight.

## 3.2 Pro Sensor Subsystem Architecture

### Biometric Sensors

1. Higher-density multi-wavelength PPG

- Wavelengths:

  - Green.

  - Red.

  - Infrared.

- Sensor configuration:

  - Three micro-LED emitter sites.

  - Four photodiode sites.

  - Multiple spatial paths for different finger geometries.

- Use cases:

  - HR.

  - HRV.

  - Resting HR.

  - SpO₂ trend.

  - Sleep pulse dynamics.

  - Pulse morphology.

2. Multi-point temperature

- Type: three miniature NTC thermistors or silicon temperature sensors.

- Placement:

  - Palmar-side primary skin contact.

  - Lateral secondary contact.

  - Internal ambient/board compensation point.

- Use cases:

  - Temperature baseline.

  - Circadian temperature trend.

  - Illness/recovery trend.

Thermal drift compensation.

3. Advanced motion sensor

- Type: ultra-low-noise 6-axis IMU.

- Optional sensor-fusion engine in package.

- Use cases:

  - Sleep movement.

  - Activity classification.

  - Motion-artifact rejection.

4. EDA / skin conductance

- Type: two low-profile stainless/titanium inner electrodes.

- Placement: two flush inner arcs separated by insulating liner.

- Use cases:

  - Stress arousal.

  - Recovery trends.

5. Capacitive fit and skin-contact sensing

- Same role as Active, with more electrodes for fit mapping.

6. Optional haptic actuator

- Type: ultra-thin rectangular or Z-axis linear resonant actuator (LRA) tuned to ~230–250 Hz; or a thin piezo bending element as a slim alternative.

- Orientation: force axis radial (toward the finger pad) for direct skin coupling, with no raised bump.

- Acoustic: low-noise unit selected to avoid an audible click, preserving silent operation.

- Use cases:

  - Fit warning.

  - Pro experimental feedback.

  - Vibrating wake-up / silent alarm.

  - Gentle haptic notifications.

## 3.3 Pro Optical Layout

Target: fit the optical system within ≤2.35 mm radial wall thickness, with no raised internal bumps.

### Layout

- Three side-firing multi-wavelength LED sites

  - One primary site at palmar contact zone.

  - Two secondary offset sites.

- Four low-profile photodiodes

  - Two large-area primary photodiodes.

  - Two smaller redundancy photodiodes.

- Twelve selectable optical paths

  - Firmware ranks channels by signal amplitude, perfusion quality, motion artifact, ambient leakage, and baseline stability.

### Optical Construction

- Flush sapphire windows with anti-smudge coating.

- Molded black optical baffles.

- Optical resin index-matched to window.

- No raised domes.

- Inner liner is gently concave to distribute pressure and keep optical windows flush against skin.

### Pro Rationale

Pro adds optical redundancy because the product promise is deeper all-day wellness and sleep insight. Instead of making Pro “more accurate at the basics,” the additional channels improve data availability, confidence scoring, and artifact rejection in difficult conditions.

## 3.4 Pro Mechanical and Industrial Design

### Target Dimensions

- Width: 5.8–6.4 mm

- Radial wall thickness: 2.05–2.35 mm

- Mass: 3.5–4.5 g

- Outer profile: premium rounded or refined concave band

- Sensor protrusions: none

### Materials

Exterior

- Polished titanium.

- Gold PVD.

- Rose-gold PVD.

- Scratch-resistant ceramic premium edition.

Interior

- Hypoallergenic ceramic-polymer liner.

- Flush conductive EDA electrodes.

- Sapphire optical windows.

- Thermally conductive micro-islands under temperature sensors.

### Water Resistance

- Target: IP68

- Pro prioritizes premium finishing and long-term comfort over high-pressure wash resistance.

- EDA electrodes are flush-sealed and electrically isolated from the shell.

### Sensor-Skin Contact

- Continuous inner comfort liner.

- No localized hard bumps.

- Slightly broader width than Active spreads pressure during sleep.

- Micro-compliant optical zones allow stable contact even during overnight swelling.

## 3.5 Pro Electronics and Processing Unit

### MCU

- ARM Cortex-M33 or M4F-class MCU with DSP/FPU.

- Larger memory than Active.

- Secure enclave or hardware crypto block.

- Sensor-hub low-power mode for overnight monitoring.

### AFE

- Dual optical AFE architecture or one higher-channel AFE plus external EDA AFE.

- More LED-current flexibility than Active.

- Lower-noise photodiode paths.

- Supports simultaneous or rapid-multiplex PPG channels.

### Battery

- Curved thin-film lithium cell.

- Capacity: 25 mAh

- More aggressive power scheduling than Active because Pro has more sensors.

### Wireless

- BLE primary.

- Local storage buffers data when away from phone, matching the product requirement for stronger sync behavior and reduced data loss.

### Haptic Driver

- Closed-loop LRA haptic driver with auto-resonance tracking, automatic overdrive and braking, low-voltage operation, and I²C control; or a high-voltage boost driver if the piezo option is used.

- Drives the actuator at its resonant frequency for maximum vibration output at minimum power.

## 3.6 Pro Firmware and Signal Processing

### Normal Mode

- PPG: periodic sampling, with higher-resolution windows during sleep.

- Temperature: multi-point sampling for stable trend.

- EDA: low-duty stress sampling.

- IMU: low-rate sleep/activity tracking.

- Fit sensing: more frequent during sleep onset and wake windows.

### Smart-Wake / Vibrating Alarm Mode

- Uses on-ring sleep-stage detection to identify a light-sleep moment within a user-set wake window (~30 min before target time) and triggers the haptic alarm then, reducing sleep inertia (morning grogginess).

- Wake pattern: short escalating pulses (gentle to strong) over ~30–90 s at the actuator’s resonance band (~230–250 Hz), not a constant buzz.

- Stop condition: movement-confirmed via the IMU (anti-snooze); never wakes earlier than the window.

- Calm Mode alignment: escalation and light-sleep timing keep waking gentle and non-startling; silent operation avoids disturbing a sleeping partner.

- Power: a wake event is a brief, once-daily pulse train and a negligible draw on the 25 mAh battery; the steady cost is the always-on sleep-stage detection already running on Pro.

# 4. Manufacturing and Calibration Approach

## 4.1 Assembly Flow

- Machine or mold outer shell.

- Prepare inner carrier with optical baffles.

- Attach flex PCB to curved assembly fixture.

- Place optical emitters, photodiodes, AFE, MCU, IMU, PMIC, wireless, memory.

- Place haptic actuator and haptic driver on the flex assembly, oriented for radial (toward-skin) vibration, within the potting/waterproof envelope.

- Attach curved battery.

- Install optical windows.

- Overmold or pot electronics.

- Final inner liner application.

- Laser mark internal product ID.

- Final calibration and waterproof testing.

The background sensor-analysis source emphasizes that miniaturization, curved flex-PCB routing, packaging, power management, and calibration are central to high-quality ring performance, not just the purchase of individual components.

## 4.2 Critical Tolerances

### Optical

- LED-to-window alignment: ±50 µm.

- Photodiode-to-window alignment: ±50 µm.

- Optical baffle height tolerance: ±30 µm.

- Window thickness: ±20 µm.

- Optical resin voids: none allowed in active path.

- Dark leakage test: required for every unit.

### Battery and Mechanical

- Battery bend radius controlled by size.

- No compression points over lithium cell.

- Potting shrinkage must not distort optical windows.

- Inner diameter tolerance: tight enough for fit accuracy, loose enough for comfort.

## 4.3 Per-Unit Calibration

### Optical Calibration

- Dark current.

- Ambient-light rejection.

- LED current sweep.

- Photodiode gain table.

- Optical crosstalk map.

- Skin-phantom reflectance test.

- Motion-artifact test on rotating finger phantom.

### Temperature Calibration

- Two-point thermal calibration.

- Thermal lag measurement.

- Board-heating compensation table.

- Pro: multi-point temperature consistency check.

### Motion Calibration

- Six-position accelerometer calibration.

- Gyro zero-rate calibration.

- Vibration/noise floor test.

- Activity-motion fixture test.

### EDA Calibration, Pro Only

- Electrode impedance test.

- Dry-contact detection.

- Known-resistance ladder calibration.

- Sweat/salt contamination check.

### Haptic Calibration (Pro Only)

- Measure actuator resonant frequency and set the driver’s auto-resonance/drive parameters per unit.

- Verify vibration output (acceleration) meets the wake threshold.

- Acoustic check: confirm no objectionable audible click.

## 4.4 Quality Control Tests

Every production unit must pass:

- Waterproof test.

- Sweat corrosion test.

- Optical signal-quality test.

- IMU test.

- Battery capacity and reserve test.

- Haptic output and resonance test (Pro): unit must produce the specified vibration force at resonance and pass the quiet-operation (no audible click) check.

# 5. Performance Targets and Technical Risks

## 5.1 Active Performance Targets

| Category | Target |
| --- | --- |
| Width | 5.2–5.8 mm |
| Radial thickness | 1.80–2.05 mm |
| Mass | 3.0–4.0 g |
| Water resistance | IP69K |
| Battery | 15 mAh curved thin-film |
| Normal battery life | 4–5 days guaranteed with |

## 5.2 Pro Performance Targets

| Category | Target |
| --- | --- |
| Width | 5.8–6.4 mm |
| Radial thickness | 2.05–2.35 mm |
| Mass | 3.5–4.5 g |
| Water resistance | IP68 |
| Battery | 25 mAh curved thin-film |
| Normal battery life | 4–5 days guaranteed with full Pro sensing profile |

## 5.3 Key Technical Risks and Mitigations

### Risk 1: Pro sensor complexity reduces battery life

Problem: More sensors can reduce runtime.

Mitigation:

- Sensor hub controls low-duty sensing.

- Pro uses adaptive sampling.

- EDA and redundant PPG channels activate mainly when context justifies.

# 6. Final Engineering Recommendation

Build both versions on the same validated core platform, because the supplied Product Idea explicitly requires shared baseline accuracy and warns against making Active intentionally weaker.

Active should be the tougher, thinner, motion-first ring:

- 15 mAh battery.

- IP69K.

- Two-emitter/two-photodiode PPG.

- Fast temperature.

- Motion-optimized IMU.

Pro should be the premium, sleep/resting-accuracy, advanced-insight ring:

- 25 mAh battery.

- IP68.

- Higher-density PPG.

- Multi-point temperature.

- EDA.

- Larger memory.

- Optional haptics.

The central breakthrough is not simply adding more sensors. The breakthrough is packaging optics, adaptive signal processing, and power management inside a ring that still looks like a high-end fashion band, with no visible sensor bumps.